TABLE OF CONTENTS

1	Introduction	1

1.1	Document overview	1

1.2	References	1

1.2.1	Project References	1

1.2.2	Standard and regulatory References	1

1.2.3	Definitions	2

1.2.4	Conventions	2

2	Responsibilities	2

3	Risk management process	3

3.1	Context establishment	4

3.2	Risk assessment	5

3.2.1	Identification	5

3.2.2	Analysis	6

3.2.3	Evaluation	7

3.3	Risk treatment	7

3.4	Risk acceptance	7

3.5	Risk communication	8

3.6	Post-production monitoring and review	8

4	Ranking system for security risk analysis	8

4.1	Probability of occurrence	8

4.2	Severity	9

4.3	Other criteria	9

4.4	Risk Priority Number	9

5	Relationships with other risk management processes	10

5.1	Ranking system of safety risks when security breach is the hazardous phenomenon	10

5.2	Communication with safety risk management team	11

5.3	Communication with usability engineering team	11

# Introduction

## Document overview

Describe the purpose of the document

This document contains the security risk management plan for XXX device. It covers the management all security-related risks during the lifecycle of the device, in design and development, and in maintenance. It also contains the provisions about relationships with the ISO 14971 safety risk management process, and IEC 62366-1 usability engineering process.

## References

### Project References

| # | Document Identifier | Document Title |
| --- | --- | --- |
| [R1] | ID | Add your documents references. One line per document |

### Standard and regulatory References

| # | Document Identifier | Document Title |
| --- | --- | --- |
| [STD1] |  | Add your documents references. One line per document |

Add the standard references to the table above. It may include ISO 14971, ISO 13485, IEC 62304, IEC 81001-5-1, AAMI SW96, AAMI TIR 57, , AAMI TIR 97, UL 2900-1, UL 2900-2-1, but also IEC 62443-x, ISO 2700x.

Add FDA guidances on premarket, post-market and OTSS, have a look at guidances on FDA website : https://www.fda.gov/medical-devices/digital-health-center-excellence/cybersecurity.

### Definitions

Information security risk: potential that a given threat will exploit vulnerabilities of an asset or group of assets, and thereby cause harm to the organization.

You may add here other definitions like Asset, Threat, Vulnerability found in IEC 81001-5-1.

### Conventions

Security risk and Cybersecurity risk are used indifferently in this document as synonyms of Information security risk according to the definition above.

# Responsibilities

Describe the organization of the team responsible for risk management. You may add an organization chart or add a reference to your project management plan, where the organization of the project should be already described.

Either describe the responsibilities in plain text:

The security risk manager is in charge of the security risk. The Quality and regulatory manager is in charge of the relationships with the ISO 14971 safety risk management process.

Or use a table

| Person | Responsibility |
| --- | --- |
| Security Officer | Overall security risk management process responsibility |
| Security Advisor | Independent reviewer of the security risk management process |
| Project Manager | Security risk management process responsibility during design |
| Hardware design manager | Hardware security risks |
| System design manager | System / network security risks |
| Software design manager | Software security risks |
| Quality Manager | Independent review of Risk Management File |
| Quality Manager | Interface with the safety risk management process |
| Deployment Manager | Interface with IT services in healthcare center |
| Customer Service Manager | Interface with end-users in healthcare center |

You may also explain here how adverse events are escalated in the organization. Eg: Adverse events or risk of adverse events detected throughout the organization shall be immediately reported to the security officer.

# Risk management process

The process described below is inspired form the risk management process presented in ISO 27005 (which stems from risk management process presented in ISO 31000), with arrangements for compliance to IEC 81001-5-1 and ISO 14971.

The risk management process is an iterative process allowing to increase the depth and details of risk assessment at each iteration.

As an alternative to the diagram below, you may also take the security risk management process diagram presented in AAMI SW96 or AAMI TIR57.

## Context establishment and threat modelling

These two steps (1 & 2 below) allow to identify assets and threats.

Establishing the context consists in defining the scope and boundaries, and describing the following items, as appropriate:

- The medical device (hardware, software, network…)

- Its accessories,

- Its environment

  - Operating room

  - Patient room

  - At home

- Other connected devices,

  - Medical devices,

  - Non-medical devices,

  - Cloud servers

- The processes involved in the lifecycle of the device:

  - Internal processes,

  - Outsourced processes and

  - Client/user processes,

- The users and user profiles

  - Client users,

  - Users of the manufacturer (e.g. customer support)

- The level of education of users,

- The use cases associated to the users and user profiles,

- The types of data handled in the device

  - Medical and personal data,

  - Data from sensors,

  - Configuration data,

  - Logs

- The hardware network interfaces

  - Bluetooth,

  - Wifi, Zigbee,

  - RJ45,

- The software network interfaces and protocols

  - HTTP, TCP, UDP,

  - SOAP, REST

  - Network Ports

- The data input/output streams

  - With connected devices,

  - Through removable media,

  - Internally between sub-systems of the device,

- The COST/SOUP used in software

  - Maintained software

  - Supported software (see IEC 81001-5-1)

- The constraints affecting the device

  - On the device,

  - On manufacturer processes,

  - On user processes, E.g. end-users generally don’t accept to un-scrub and scrub for IT security reasons,

  - Regulatory requirements: GDPR, HIPAA…

- Constraints regarding emergency access to the device for patient safety, bypassing security measures.

All items listed above in 2nd level bullets are practical examples, other cases may be applicable to your device.

Assets

An asset is anything of value for the manufacturer or for the end-user.

Assets can be:

- Hardware,

- Software,

- Data,

- SW Services in interface with the device,

- SW development platform,

- Other supporting services,

- Other tangible or intangible assets.

Examples:

- The medical device itself,

- Its accessories,

- Devices connected to the medical device,

- Cloud servers,

- Patient data,

- Diagnosis or Treatment data,

- Other data.

Knowing the assets contributes to identify the magnitude of the consequences of an attack. Eg: the consequences won’t have the same order of magnitude if the device is a simple object, a cloud server, or a $1M x-ray scanner.

You may reference here the list of assets in annex B of AAMI TIR 57 or definition 3.4 of IEC 81001-5-1 as a help to identify assets.

Threat actors

A threat is anything (human, phenomenon, accidental, deliberate) susceptible to result in a damage on one or more assets.

Examples:

- Script kiddies (eg.: the patient’s nephew, the teenager in their room)

- Academic researchers

- Criminal organizations, advanced persistent threats,

- Inexperienced users (the patient, a caregiver, a dismounted employee),

- Natural events.

You may reference here the list of threats in annex B of AAMI TIR 57 as a help to identify threats.

This context allows to identify the threats, documented in the treat modelling. The context and threat modelling can be described in data flow diagrams (DFD), use cases, hardware architecture, system architecture, software architecture, as appropriate.

Some information may already be documented in the IEC 62336-1 usability engineering file and shall be referenced by the threat modelling documentation. E.g. Users, user profiles, use cases.

The constraint of emergency access to the device, bypassing security measures, shall be described in the context.

Additionally, the manufacturer shall provide justification for any exclusion from the scope.

See the threat modelling playbook. It’s free and will help a lot for this task: https://mdic.org/resource/playbook-for-threat-modeling-medical-devices/#download_form

Note: this playbook has been financed by the FDA.

## Risk assessment

Based on the threat model documented in the previous step, the risks shall be identified, analyzed and evaluated according to the objectives of the risk management process and the ranking system for security risk analysis.

A preliminary risk assessment, together with the threat modelling, may be documented at the beginning of the design project, or before design (e.g. in feasibility / mock-up etc.) and reviewed as part of design input data.

### Identification

(steps 3, 4 and 5 below)

Existing controls.

Existing or planned controls should be identified. They can be identified inside the manufacturer’s organization or in the target environments or other environments (healthcare provider, datacenter service supplier…).

Existing or planned controls should be assessed to determine if they are effective, sufficient and justified.

Weaknesses and Vulnerabilities

Weaknesses and Vulnerabilities shall be identified.

Weaknesses may come from design decisions, anticipated misuse of the device, organizations, processes, etc.

Vulnerabilities can be found in catalogs like the Common Vulnerabilities and Exposures List, the National Health ISAC and release notes of SOUP /COTS editors.

You may reference here the annex B of AAMI TIR 57 as a help to identify vulnerabilities.

Consequences

The consequences that loss of confidentiality, integrity and availability may have on the assets shall be identified. The consequences on patient safety shall also be identified.

You may reference here the section Impact Assessment of annex B of AAMI TIR 57 as a help to identify consequences.

Records

Assets, threats, vulnerabilities, existing controls and consequences shall be recorded in the security risk assessment report.

### Analysis

The risk analysis is performed with the use of the ranking system described in section 4 of this document, and with the data collected in the previous steps:

- The identification of the consequences on assets allow to determine the impact magnitude of the risk,

- The identification of threats, vulnerabilities, and incident scenarios allow to determine the likelihood of a risk.

Records

For each incident scenario: probability of occurrence, severity, other criteria, and RPN are recorded in the risk assessment report.

### Evaluation

The RPN are compared against risk acceptance criteria described in section 4. Legal and regulatory requirements shall be included in the evaluation of the acceptance of risks.

Records

For each incident scenario: risk acceptance, and decision.

## Risk treatment

The term “risk treatment” is used in ISO 27005. Note that the meaning of “risk treatment” is broader than “risk control” or “risk mitigation” commonly used in medical device industry.

For information, risk treatment options are:

- Mitigate (thus risk control)

  - Take action to reduce the likelihood that the threat will materialize.

- Eliminate (kind of inherent security by design)

  - Simply remove the feature or component that is causing the threat.

- Transfer (kind of information to users)

  - Shift responsibility to another entity such as the customer or the integrator of the device.

- Accept (thus, need of information to users)

  - Do not mitigate, eliminate, or transfer the risk because none of the above options are acceptable given business requirements or constraints.

Note that risk transfer and acceptance shall be acceptable from a safety (ISO 14971) point of view. These to kind of actions couldn’t be possible for security risks with a safety impact.

Risk controls shall be sought in this order of precedence (see AAMI SW96):

| Risk Control | Example |
| --- | --- |
| Inherent security by design; | Design change in hardware or software, documented in a HRS or SRS |
| Inherent security by manufacturing process; | Process step making use of state-of-the-art security feature. E.g.: use of digital signature to verify the authenticity and integrity of a software installation bundle |
| Protective threat mitigation measures added into the medical device; | Hardware or software feature added to the device. E.g.: delivering the device with a firewall |
| Protective threat mitigation measures added to the manufacturing process; | Process step making use of less effective security feature (Compared to inherent security by manufacturing process). E.g.: ensuring software authenticity by organizational measure, testing software after install |
| Security risk disclosure (i.e., providing needed security information to users) and/or security risk transfer to users; and | Secure operations guidelines, Warnings about secure use of the device E.g.: reminder in the Secure operations guidelines to activate security windows updates by default |
| Where appropriate, security training to users. | Training to users is deemed poorly effective |

Risk control should be carried out to decrease the RPN of the residual risk As Low As Reasonably Possible (ALARP), with economic considerations. If the risk has an impact on patient safety, the risk treatment (risk control) shall be carried out with acceptance criteria related to patient safety.

The security vulnerabilities or impacts on patient safety arising from risk treatment shall also be assessed.

Records

For each incident scenario: the risk treatment plan.

## Risk acceptance

The reevaluation of the risk is performed with the use of the ranking system described in section 4 of this document, and with the risk treatment plan.

A security risk may be deemed accepted even if it doesn’t match the risk acceptance criteria of section 4, with a justification based on the risk/benefit ratio, to override the acceptance criteria.

Records

For each incident scenario: the risk acceptance and justification as appropriate.

## Risk communication

The risk assessment report is communicated internally to the design team. The risk assessment report shall be an agenda item of design reviews, and validation reviews.

Parts of the risk assessment report may also be communicated externally to service providers or suppliers, as appropriate. When a supplier requires knowledge of the risk assessment report, it should be determined if this supplier shall be placed in the list of critical suppliers.

Add also the need to communication to ISAC ISAO, per FDA guidance on cybersecurity, if needed.

## Post-production monitoring and review

The Risk Management File is systematically reviewed and updated, especially when:

- The product is modified,

- The context changes,

- Assets, threats, threat actors, vulnerabilities change,

- Analysis of data for production processes triggers a reevaluation,

- Analysis of data of post marketing surveillance triggers a reevaluation (internal defects, customer requests, maintenance, vigilance bulletins, security incidents, field information from any source see AAMS SW96 for source of information).

The review of post-marketing data is performed quarterly / annually / other period. Reviews and updates to any risk will be documented, approved, and included within the Risk Management File.

The review includes an evaluation of the relevance of the ranking system and the need to update it upon business or regulatory context.

# POSSIBILITY 1: Ranking system for security risk analysis based on CVSS

According to UL 2900-1, the risk criteria shall be based on CVSS or derived from CVSS

When claiming compliance to UL 2900, a possibility is to use the probability based on CVSS:

See https://www.first.org/cvss/calculator/3.1

P = CVSS base score [0 ; 10]

| Ranking | Definition | Comment |
| --- | --- | --- |
| Critical | 9 ≤ CVSS base score ≤ 10 | Risk is unacceptable |
| High | 7 ≤ CVSS base score < 9 | Risk is unacceptable |
| Medium | 4 ≤ CVSS base score < 7 | Risk is tolerable |
| Low | 0 < CVSS base score < 4 | Risk is acceptable |
| None | 0 = CVSS base score | No risk |

Remark, it may be possible to tailor the acceptability to the SL-C (see IEC TR 60601-4-5) of the software:

| Ranking | Definition | SL-C 1 | SL-C 3 |
| --- | --- | --- | --- |
| Critical | 9 ≤ CVSS base score ≤ 10 | Risk is unacceptable | Risk is unacceptable |
| High | 7 ≤ CVSS base score < 9 | Risk is tolerable | Risk is unacceptable |
| Medium | 4 ≤ CVSS base score < 7 | Risk is acceptable | Risk is tolerable |
| Low | 0 < CVSS base score < 4 | Risk is acceptable | Risk is acceptable |
| None | 0 = CVSS base score | No risk | No risk |

# POSSIBILITY 2 Ranking system for security risk analysis

This second possibility may be more convenient for medical device manufacturer teams, as it is closer to what they know with safety risk management

This section describes how the risk priority number is deduced from the characteristics of the risk:

- Likelihood of occurrence,

- Severity of consequences

- Other criteria (if any).

Describe in sub sections how you quantify your criteria, like these:

## Likelihood of occurrence

You can use quantitative or qualitative criteria. Here is an example.

| Ranking | Definition | Likelihood (L) |
| --- | --- | --- |
| 5 | Often occurs, once a week | Frequent (very high likelihood) |
| 4 | Could easily happen, once a month | Probable (high likelihood) |
| 3 | Could happen or known to happen, once a year | Occasional (moderate likelihood) |
| 2 | Hasn’t happened yet but could, once in device lifetime | Unlikely (low likelihood) |
| 1 | Conceivable but only in extreme circumstances, less that once in device lifetime | Very Unlikely (very low likelihood) |

Here is another example form the NIST SP 800-30 Rev1 Guide for Conducting Risk Assessments, where the likelihood is split in two elements:

And:

## Severity

Use quantitative or qualitative criteria. Here is an example. You may add explicit links with consequences on HIPAA or GDPR compliance.

Note: the definitions in the below table always contain a wording based on consequences on Confidentiality, Integrity and Availability, plus economic consequence.

| Ranking | Definition | Severity |
| --- | --- | --- |
| 5 | Significant loss of confidentiality or data/system integrity or system availability. Threatens financial health of manufacturer. | Catastrophic |
| 4 | Some loss of confidentiality or or data/system integrity or system availability. Some consequences on financial health of manufacturer. | Critical |
| 3 | Small, isolated loss of confidentiality or or data/system integrity or system availability. Small consequences on financial health of manufacturer. | Moderate |
| 2 | Isolated, limited leak of non-sensitive data. Isolated loss of data/system integrity or system availability May incur costs to recover leaked data. | Minor |
| 1 | Recoverable data loss, system availability/integrity loss, Non-sensitive data leak. Negligible cost. | Negligible |

## Other criteria

Add other criteria, or remove this section

## Risk Priority Number

A rule of your choice, like.

Risk priority number = 	criterion 1

x criterion 2

…

x criterion n

Example of cross-table of RPN with two criteria:

|  | CROSS TABLE OF RISK PRIORITY NUMBER |
| --- | --- |
|  | Negligible 1 | Minor 2 | Moderate 3 | Critical 4 | Catastrophic 5 |
| Frequent 5 | 5 tolerable | 10 tolerable | 15 un-acceptable | 20 un-acceptable | 25 un-acceptable |
| Probable 4 | 4 acceptable | 8 tolerable | 12 un-acceptable | 16 un-acceptable | 20 un-acceptable |
| Occasional 3 | 3 acceptable | 6 tolerable | 9 tolerable | 12 un-acceptable | 15 un-acceptable |
| Unlikely 2 | 2 acceptable | 4 acceptable | 6 tolerable | 8 tolerable | 10 tolerable |
| Very Unlikely 1 | 1 acceptable | 2 acceptable | 3 acceptable | 4 acceptable | 5 tolerable |

# Relationships with other risk management processes

The safety risk management team may be separate from the security risk management team and the usability engineering team. Or they may be in a single team. In any cases, communication is required between safety, usability, and security teams.

See also figure 4 of AAMI TIR 57

## Ranking system of safety risks when security breach is the hazardous phenomenon

According to Annex F of ISO 24971:2020, in case of security risks with impact on safety, the probability of risks can be split into:

- P1 = Probability that a vulnerability can be exploited by the threats (the hazardous situation),

- P2 = Probability that the hazardous situation (a vulnerability can be exploited by the threats) leads to a harm.

POSSIBILITY 1: Ranking system for security risk analysis based on CVSS Exploitability Metrics

P1 of safety risk can be estimated with the CSSS 3.1 Exploitability metrics [0.12 ; 3.89]

| Ranking | Definition | Probability (P1) |
| --- | --- | --- |
| 5 | 3.13 ≤ Exploitability ≤ 3.89 | Frequent (very high probability) |
| 4 | 2.38 ≤ Exploitability < 3.13 | Probable (high probability) |
| 3 | 1.63 ≤ Exploitability < 2.38 | Occasional (moderate probability) |
| 2 | 0.87 ≤ Exploitability < 1.63 | Unlikely (low probability) |
| 1 | 0.12 ≤ Exploitability < 0.87 | Very Unlikely (very low probability) |

POSSIBILITY 2: Ranking system for security risk analysis based on likelihood

P1 of safety risk can be estimated as equal to the likelihood of the security risk as defined above.

| Ranking | Definition | Probability (P1) |
| --- | --- | --- |
| 5 | Very High Likelihood | Frequent (very high probability) |
| 4 | High Likelihood | Probable (high probability) |
| 3 | Moderate Likelihood | Occasional (moderate probability) |
| 2 | Low Likelihood | Unlikely (low probability) |
| 1 | Very Low Likelihood | Very Unlikely (very low probability) |

P2 is estimated with the scale:

| Ranking | Definition | Probability (P2) |
| --- | --- | --- |
| 5 | Frequent | Attack systematically or very often leads to harm |
| 4 | Probable | Attack could easily lead to harm in normal condition of use |
| 3 | Occasional | Attack could lead to harm in particular condition of use |
| 2 | Remote | There’s a chance that the attack will lead to harm in device lifetime |
| 1 | Improbable | There’s no chance that the attack will lead to harm in device lifetime |

Then the total probability of a safety risk can be estimated as P = P1 x P2.

| Ranking | Definition | Description |
| --- | --- | --- |
| 5 | Frequent | 20 ≤ P1xP2 ≤ 25 |
| 4 | Probable | 15 ≤ P1xP2 < 20 |
| 3 | Occasional | 10 ≤ P1xP2 < 15 |
| 2 | Remote | 5 ≤ P1xP2 < 10 |
| 1 | Improbable | 0 < P1xP2 ≤ 5 |

POSSIBILITY 3: If you don’t have discrete P1 / P2 in your safety risk probability modelling

Just add:

The probability of safety risk where the hazardous situation is a security breach is based on the likelihood of the security breach and the probability of the breach to give rise to a harm.

## Communication with safety risk management team

Communication with the safety risk management team shall be performed in a timely manner when:

- A security risk has a potential impact on safety,

- Security controls may affect safety,

- A security incident may affect safety.

In case of security incident, which may affect safety, the information shall be immediately reported to the person in charge of medical device reporting / materiovigilance.

## Communication with usability engineering team

Communication with the usability engineering team shall be performed in a timely manner when:

- A security risk has a potential impact on usability,

- A security control may affect or have an impact on usability.

When a security control has an impact on usability, the assessment of the effectiveness of the security control shall take account of results of usability assessment. Eg. the use scenario where the security control appears may be included in the summative evaluation.

# Text boxes

More templates to download on the:

Templates Repository for Software Development Process (click here)

Or paste the link below in your browser address bar:

http://blog.cm-dm.com/pages/Software-Development-Process-templates

This work is licensed under the:

Creative Commons Attribution-NonCommercial-NoDerivs 3.0 France License: http://creativecommons.org/licenses/by-nc-nd/3.0/fr/

Waiver:

You can freely download and fill the templates of blog.cm-dm.com, to produce technical documentation. The documents produced by filling the templates are outside the scope of the license. However, the modification of templates to produce new templates is in the scope of the license and is not allowed by this license.

To be compliant with the license, I suggest you to keep the following sentence at least once in the templates you store, or use, or distribute:

This Template is the property of Cyrille Michaud License terms: see http://blog.cm-dm.com/post/2011/11/04/License

Who am I? See my linkedin profile:

http://fr.linkedin.com/pub/cyrille-michaud/0/75/8b5

You can remove this first page when you’ve read it and acknowledged it!

Thank-you for downloading the

Security Risk Management Plan Template!