TABLE OF CONTENTS

1	INTRODUCTION	2

1.1	Document overview	2

1.2	Abbreviations and Glossary	2

1.2.1	Abbreviations	2

1.2.2	Glossary	2

1.3	References	2

1.3.1	Project References	2

1.3.2	Standard and regulatory References	2

1.4	Conventions	2

2	REQUIREMENTS	4

2.1	States	4

2.2	Performance	4

2.3	Safety, security, and privacy protection	4

2.4	User maintenance	5

2.5	Usability and human-factors engineering	5

2.5.1	Man machine interface layout	5

2.5.2	Help	5

2.6	System environment	6

2.7	External interfaces	6

2.7.1	Hardware interfaces	6

2.7.2	Network interfaces	6

2.7.3	Data exchange	6

2.8	Resources	6

2.8.1	Hardware resources	6

2.8.2	Software resources	6

2.9	Internal data	6

2.10	Adaptation	7

2.11	Verification	7

2.12	Personnel and training	7

2.13	Packaging and installation	7

3	VERIFICATION METHODS	8

4	REQUIREMENTS TRACEABILITY	10

5	CRITICAL REQUIREMENTS	11

# INTRODUCTION

## Document overview

This document presents the software requirements specifications of XXX software development project.

It describes:

- Requirements of functionalities, performances, interfaces, environment …

- Tests principles and definitions of validation methods of requirements,

- The compliance of requirements to customer needs,

- The relative importance and precedence of requirements

## Abbreviations and Glossary

### Abbreviations

Add here abbreviations

### Glossary

Add here words definitions

## References

### Project References

| # | Document Identifier | Document Title |
| --- | --- | --- |
| [R1] | ID | Add your documents references. One line per document |

### Standard and regulatory References

| # | Document Identifier | Document Title |
| --- | --- | --- |
| [STD1] |  | Add your documents references. One line per document |

## Conventions

Requirements listed in this document are constructed according to the following structure:

Requirement Id

Requirement title

Requirement description

Requirement version

Example:

SRS-XXX-000

Title of XXX-000 requirement

Description of XXX-000 requirement

Version of XXX-000

# REQUIREMENTS

Note : have a look at  http://en.wikipedia.org/wiki/Requirement, article in wikipedia. It’s well written and the links at the bottom are useful.

## States

FOO software works in three states:

- Starting: the software loads its components;

- In use: all the functionalities of the software are available to the users;

- Stopping: the software is being stopped.

- Maintenance: the software is in maintenance mode

- And so on …

Add a diagram with states and transitions if necessary

## Performance

This is the core of your SRS. It contains the purpose of your software expressed in technical requirements

SRS-XXX-010 SAMPLE

Sample requirement about a function

FOO software shall compute the zzz parameters with the a, , c and d input parameter, with the use of the XXX algorithm.

V1.0

SRS-XXX-020 SAMPLE

Sample requirement about a function

FOO software shall save the result of computations in boo-bar format.

V1.0

## Safety, security, and privacy protection

This section is about software features like confidentiality, integrity control, reliability, and availability. See CyberSecurity requirements of FDA and HIPAA requirements if necessary

SRS-XXX-030 SAMPLE

Patient data

XXX ensures that the displayed patient data are the same as read in the input files.

The patient’s data are:

- Name,

- Date of birth,

V1.0

## User maintenance

SRS-XXX-040 SAMPLE

Application logs

XXX generates a log file containing:

- The state of the application and the steps performed to reach that state,

- The possible error logs, if any.

V1.0

## Usability and human-factors engineering

The requirements here may have traceability with result of 62366 standard implementation

### Man machine interface layout

The layout of XXX is ….

Instead of a dozen of text requirements, a mock-up of the software GUI is very appreciated

Add only requirements for which a description of layout/behaviour is necessary and/or requested by a user.

SRS-XXX-050 SAMPLE

Menu items and other widgets

XXX software has the following items:

- Menu file ...

- Widgets in the main window (slider, button, radiobutton, textfield).

V1.0

### Help

The user guide is always very important for medical devices. It may be online, in this case add requirements here about the online help ….

SRS-XXX-060 SAMPLE

Online user guide

XXX contains an online user guide

V1.0

An about window is a good way to identify software version….

SRS-XXX-070 SAMPLE

About XXX

XXX shall display an “About…” window. This window displays the current version of the application.

V1.0

## System environment

If software is integrated in a specific system, describe briefly the system and add specific requirements to which your software shall comply

Warning : for PEMS/Electro-medical Devices with a big system architecture, a system architecture document is necessary to describe the system/software architecture.

## External interfaces

This section describes hardware and software interfaces of the software in the system

### Hardware interfaces

For PEMS/Electro-medical Devices, add requirements about integration of software and hardware.

### Network interfaces

Also add here communication and networks stuff, like IP, wireless, Bluetooth …

### Data exchange

If XXX software is in interface with other software, describe here the requirements on data exchanges.

## Resources

### Hardware resources

SRS-XXX-080 SAMPLE

Hardware configuration

XXX shall run with the expected response times on a PC with the following minimal configuration:

- 2 Go RAM

- ...

V1.0

### Software resources

SRS-XXX-090 SAMPLE

Software configuration

XXX runs in the following software environment:

- (describe OS version),

V1.0

## Internal data

If specific requirements for internal data, like databases, binary files, xml …

## Adaptation

If specific requirements adaptability of configuration of software

## Verification

Special functions to test the software, if necessary. For example a hidden function to activate a log file during beta tests

## Personnel and training

Requirements about the capabilities/knowledge of users, the training the shall have before using software

SRS-XXX-USR-010 SAMPLE

E-learning

XXX is delivered with e-learning module.

V1.0

## Packaging and installation

SRS-XXX-PAK-010 SAMPLE

Packaging

XXX shall be delivered on zzz media.

V1.0

SRS-XXX-PAK-010 SAMPLE

Install-shield

XXX shall be installed with the use of an install shield.

V1.0

# VERIFICATION METHODS

The verification methods of the requirements are defined below:

- Inspection (I): control or visual verification

  - Control of the physical implementation or the installation of a component. The control verifies that the implementation or the installation of a component is compliant with the requirements of diagrams.

  - Control of the documentation describing a component. The control verifies that the documentation is compliant with the requirements.

- Analysis (A): verification based upon analytical evidences

  - Verification of a functionality, performance or technical solution of a component by analyzing the data collected by tests in real conditions, by simulation of real conditions or by a analysis report.

  - Analysis of test data or of design data is used as appropriate to verify requirements.

  - The verification is based upon analytical evidences obtained by calculations, like modeling, simulation and forecasting.

  - Analysis is used when an acceptable level of confidence cannot be established by other methods or if analysis is the most cost-effective solution.

- Demonstration (D): verification of operational characteristics, without quantitative measurement

  - Verifying a requirement by demonstration implies that the required functionality specified by a requirement is complete.

  - Demonstration is used when quantitative measurement is not required for verification of the requirements

  - Demonstration includes the control of the technical solutions specified by the non-functional requirements.

- Test (T): verification of quantitative characteristics with quantitative measurement

  - Verifying a functionality, performance or technical solution of a component by executing testing scenarios in predefined, controlled and traceable testing conditions.

  - Tests require the use of special equipment, instrumentation, simulation techniques, or the application of established principles and procedures,

  - Data produced during tests is used to evaluate quantitative results and compare them with requirements.

For each requirement of the SRS, a verification method is defined. Method is abbreviated I, A, D or T.

| Requirement ID | Requirement Title | Method |
| --- | --- | --- |
| REQ-001 | Verify that the speed is displayed in rpm | D |
| REQ-001 | Verify that the color of background is blue | I |

Note: do not mistake the two meanings of the word “test” in this document:

- The method of verification, named Test and abbreviated (T), as defined above.

- A test, or test case, is a sequence of actions to verify a requirement. Tests are defined in the software test plan.

Examples of tests methods:

Inspection:

- Verify that the color of background is blue,

- Verify that the user manual has the CE mark on its cover

- Verify that the PC has 4Gb memory

- Verify that firmware version on electronic card is 1.0.1

Demonstration

- Verify that when the user closes the window, a confirmation message appears

- Verify that the file is saved in the output directory

- Verify that the result is shown

- Verify that if a value is out of range, a warning is displayed

Analysis:

- Verify that the statistical distribution of results of xxx algorithm is a Gaussian with mean=x and stdev=y, when input data are blah blah

- Verify that the linear regression of results of xxx algorithm is a line which value is 1 on the y-axis, at zero on the x-axis,

Test:

- Verify that a file of 1Gb is processed in less than 3s

- Verify that the response time of the server is 15ms with 20 simultaneous requests

Rule of thumb for software, 80% of requirements are verified by demonstration, 15% by inspection and 5% by analysis or test methods.

# REQUIREMENTS TRACEABILITY

Add a table with traceability of technical requirements of this document with functional requirements.

Example

| SRS Req. | Req Title | Functional Req. | Req. Title |
| --- | --- | --- | --- |
| SRS-REQ-001 | Reading ECG values | FUN-REQ-00A | ECG post treatment |
| SRS-REQ-002 | Writing results | FUN-REQ-00A | ECG post treatment |

# CRITICAL REQUIREMENTS

If necessary, add a list of critical requirements, or a list of reference to requirements in previous sections.

This list may be the result of risk analysis (ISO 14971).

Examples

| Requirement ID | Requirement Title | Origin |
| --- | --- | --- |
| REQ-001 | Alarm when value out of range | Risk Analysis |
| REQ-002 | Do not open file if no patient name | Risk Analysis |
| REQ-003 | Display negative values in red color | Human factor engineering |

# Text boxes

More templates to download on the:

Templates Repository for Software Development Process (click here)

Or paste the link below in your browser address bar:

http://blog.cm-dm.com/pages/Software-Development-Process-templates

This work is licensed under the:

Creative Commons Attribution-NonCommercial-NoDerivs 3.0 France License: http://creativecommons.org/licenses/by-nc-nd/3.0/fr/

Waiver:

You can freely download and fill the templates of blog.cm-dm.com, to produce technical documentation. The documents produced by filling the templates are outside the scope of the license. However, the modification of templates to produce new templates is in the scope of the license and is not allowed by this license.

To be compliant with the license, I suggest you to keep the following sentence at least once in the templates you store, or use, or distribute:

This Template is the property of Cyrille Michaud License terms: see http://blog.cm-dm.com/post/2011/11/04/License

Who am I? See my linkedin profile:

http://fr.linkedin.com/pub/cyrille-michaud/0/75/8b5

You can remove this first page when you’ve read it and acknowledged it!

Thank-you for downloading the

Software Requirements Specifications Template!