TABLE OF CONTENTS

1	Introduction	1

1.1	Document overview	1

1.2	References	1

1.2.1	Project References	1

1.2.2	Standard and regulatory References	1

2	Risk Analysis	1

2.1	Intended use	1

2.2	End users	2

2.3	Foreseeable misuse	2

2.4	Characteristics Affecting Safety	2

2.5	Software classification	2

2.6	Risk analysis and evaluation	3

2.7	Risk traceability matrix	5

2.8	Overall assessment of residual risks	7

# Introduction

## Document overview

This document covers the security risk assessment report of XXX device, designed in XXX software development project.

It contains:

- The risk analysis,

- The risk assessment report,

- The risk traceability matrix with software requirements.

## References

### Project References

| # | Document Identifier | Document Title |
| --- | --- | --- |
| [R1] | ID | Add your documents references. One line per document |

### Standard and regulatory References

| # | Document Identifier | Document Title |
| --- | --- | --- |
| [STD1] |  | Add your documents references. One line per document |

Add the standard references to the table above. It may include ISO 14971, ISO 13485, IEC 62304, IEC 81001-5-1, AAMI TIR 57, UL 2900-1, UL 2900-2-1, but also ISO 2700x.

Add FDA guidance, have a look at guidances on FDA website : https://www.fda.gov/medical-devices/digital-health-center-excellence/cybersecurity.

# Risk Analysis

## Intended use

Paste here intended use

## Context of risk assessment

Describe here the context of the risk assessment, as required in the risk management plan

The context may contain the following items, as appropriate:

- The medical device (hardware, software, network…)

- Its accessories,

- Its environment

  - Operating room

  - Patient room

  - At home

- Other connected devices,

  - Medical devices,

  - Non-medical devices,

  - Cloud servers

- The processes involved in the lifecycle of the device:

  - Internal processes,

  - Outsourced processes and

  - Client/user processes,

- The users and user profiles

  - Client users,

  - Users of the manufacturer (e.g. customer support)

- The level of education of users,

- The use cases associated to the users and user profiles,

- The types of data handled in the device

  - Medical and personal data,

  - Data from sensors,

  - Configuration data,

  - Logs

- The hardware network interfaces

  - Bluetooth,

  - Wifi, Zigbee,

  - RJ45,

- The software network interfaces and protocols

  - HTTP, TCP, UDP,

  - SOAP, REST

  - Network Ports

- The data input/output streams

  - With connected devices,

  - Through removable media,

  - Internally between sub-systems of the device,

- The COST/SOUP used in software

- The constraints affecting the device

  - On the device,

  - On manufacturer processes,

  - On user processes, E.g. end-users generally don’t accept to un-scrub and scrub for IT security reasons,

  - Regulatory requirements: GDPR, HIPAA…

- Constraints regarding emergency access to the device for patient safety, bypassing security measures.

Some information may already be documented in the IEC 62336-1 usability engineering file. E.g. Users, user profiles, use cases.

### Assets

The list of assets is

- The medical device itself,

- Its accessories,

- Devices connected to the medical device,

- Cloud servers,

- Patient data,

- Device configuration data,

- Diagnosis or Treatment data,

- Private keys,

- Other data.

### Threats and threat model

Expected threats for the medical device, given its context of use:

- Script kiddies

- Academic researchers

- Criminal organizations

- Inexperienced users

- Natural events

Add a reference to the threat modelling documents.

## Security risk matrix

The matrix below contains the risk analysis table, used for the study of the security risks associated with the device.

Add here a matrix with risk analysis.

Given the variety of risk analysis methods, the matrix may have different forms. The risk analysis method shall be described in the risk management plan. See the possible approaches in B.4 of AAMI TRI 57, threat-oriented, asset/impact-oriented, vulnerability-oriented.

For the probability, you may add columns for the details of how the probability is computed:

- Attack vector,

- Attack complexity,

- Privilege required,

- User interaction,

- Scope,

- Confidentiality impact,

- Integrity impact,

- Availability impact.

| ID | ASSET | THREAT | VULNERA-BILITY | EXISTING CONTROLS | CONSE-QUENCES | CVSS Vector | CVSS | DECISION, RISK TREATMENT | R.A. M.A.* | CVSS Vector | CVSS | SAFETY RISK? ** |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Medical device | An intruder can exploit the password weakness to break into the system | Password is vulnerable for dictionary or exhaustive key attacks | Password is required but no additional provision exist | The resources within the device are prone for illegal access/ modify/ damage by the intruder | CVSS:3.1/AV:L/AC:L/PR:N/UI:N/S:U/C:H/I:H/A:H | 8.4 | Risk control: Implement password strength and expiration policy: -implement user password rules in OS -create instruction for password policy for end-users | N/A | CVSS:3.1/AV:L/AC:H/PR:H/UI:N/S:U/C:L/I:L/A:L | 3.9 | YES, potential damage in the device leading to delayed treatment, see risk #yy |
| 2 | Medical device | A malicious attack occurs, which may be a side effect of a broad attack with the MD is not specifically targeted | Well-known vulnerability in OS, for which a patch exists but exploited by attackers with low technical level | OS security patches are applied on the medical device on a regular basis | The resources within the device are prone for illegal access/ modify/ damage by the intruder | CVSS:3.1/AV:L/AC:L/PR:N/UI:N/S:U/C:H/I:H/A:H | 8.4 | Risk control: -Disable network services and protocols, which are not required for the proper use of the device - | N/A | CVSS:3.1/AV:N/AC:H/PR:H/UI:R/S:U/C:L/I:L/A:L | 3.9 | YES, potential damage in the device leading to delayed treatment, see risk #zz |
| 3 | Patient data | An intruder can read database data and unveil patient data | Database storage in plain text | None | Loss of confidentiality | CVSS:3.1/AV:A/AC:L/PR:N/UI:N/S:U/C:H/I:N/A:N | 6.5 | Risk control: -Authorize only localhost connection to the database -Cypher the column with patient data | N/A | CVSS:3.1/AV:L/AC:H/PR:N/UI:N/S:U/C:L/I:N/A:N | 2.9 | NO, no impact on, safety |
| 4 | Medical Device | An intruder can exploit remote access weakness to break into the system | Well-known vulnerability in OS, for which a patch exists but exploited by attackers with low technical level | OS security patches are applied on the medical device on a regular basis | The resources within the device are prone for illegal access/ modify/ damage by the intruder | CVSS:3.1/AV:L/AC:L/PR:N/UI:N/S:U/C:H/I:H/A:H | 8.4 | Risk control: -Authorize remote access with two factors authentication | Yes, impact on usability, leading delayed treatment, see risk #zz | CVSS:3.1/AV:L/AC:H/PR:N/UI:N/S:U/C:L/I:N/A:N | 2.9 | YES, see column RAMA |

*R.A.M.A: Risk arising from mitigation action, either security risk (then add a line for this risk in this matrix), or a safety risk (then add it to the safety risk matrix)

**The risk itself or the risk treatment has an impact on safety risk assessment?

## Risk traceability matrix

The risk traceability matrix below contains the connections between the risk analysis, software requirements and test plan.

A risk is deemed mitigated when the test status is set to PASSED in the test report.

Traceability is a central activity of software design. The best way to ensure that a risk is mitigated, is to add a requirement in the software requirement specification (SRS). The requirement will be tested by one or more tests according to the test plan. When all the tests are PASSED, we have the proof that the risk is mitigated.

Some risks may be mitigated by other elements than software requirements, for example warnings in the instruction for use. These requirement about non-software elements can nonetheless be added to the SRS. See my SRS template for some samples.

| ID | RISK | SRS REQUIREMENT ID | SRS REQUIREMENT TITLE | TEST ID | TEST TITLE | COMMENT |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Vulnerable password can be exploited by an intruder | SRS-REQ-001 | Password strength | TEST-REQ-001 | Verify password strength rules | Four requirements and four tests to mitigate the risk #1 |
| 1 |  | SRS-DOC-001 | Password instruction | TEST-DOC-001-1 | Verify that password instruction exists |  |
| 1 |  | SRS-REQ-002 | Password expiration | TEST-REQ-002-1 | Verify that password expires after xx days |  |
| 1 |  | SRS-REQ-003 | Password history | TEST-REQ-003-1 | Verify that the last 3 passwords cannot be reused |  |

Most of times, there is a one-to-many relationship between risks, mitigation requirements, and tests verifying requirements. The example above shows that 4 requirements were defined to mitigate the risk and that 4 tests are necessary to prove that the risk is mitigated.

The risk traceability matrix below contains the connections between the risk analysis and software architectural or detailed design.

| ID | RISK | SOFTWARE ELEMENT | SOFTWARE UNIT | COMMENT |
| --- | --- | --- | --- | --- |
| 1 | Weak Password | OS: password | N/A | Mitigation of risk 1 |
| 2 | Unsecured network communication | OS: network | N/A | Mitigation of risk 2 |

Most of times, there is a one-to-many relationship between risks and software elements or software units. Quote the relevant element/unit that bear or mitigates the risk, or quote all elements/units. This depends on your software architecture. SOUPs can be involved in the traceability, especially when the OS implements security features

## Overall assessment of residual risks

Write here a qualitative assessment that the overall residual risk is acceptable. The justification may be grounded on results of penetration testing after implementation of risk treatment plans.

The qualitative assessment may be also based on impact (or absence of impact) of security risks or their mitigation on safety risks or usability.

The qualitative assessment may also be based on the device risk/benefit ratio, or it may give hints for the assessment of the device risk/benefit ratio found in the clinical evaluation report.

## Risk communication

Write here to whom the risk assessment report or parts of the risk assessment report is communicated, for what purpose and how frequently.

The safety risk management team and the usability engineering team shall not be forgotten!

Risk communication also targets the accompanying documents, with relevant information on how to secure the device and inform users on residual vulnerabilities